High Threshold Cerenkov Counter Maintenance and Operations Manual

Youri Sharabian, Jefferson Lab

October 31, 2016

CONTENT

- Introduction…………………………………………………...…………….3

- Main Components and Equipment………………………………….......3

2.1 Containment Vessel……………………………………………….…..3

2.2 Combined Ellipsoidal Mirror……………………………………......4

2.3 Photomultiplier Mounting Units………………………………...….5

2.4 Light Monitoring System………………………………………….…6

2.5 Gas Supply System…………………………………………...……….7

- Standard Operations……………………………………………….………8

3.1 Alarm Handler……………………………………........................……8

3.1.1 Gas System Alarms………..…………………………………...…..8

3.1.2 High Voltage Alarm…………………………………………....…..9

3.2 Gas Supply System……………………………………………………9

3.3 High Voltage System…………………………..……………………..9

3.3.1 High Voltage Power Supply……………………………………….9

3.3.2 High Voltage Controls…………………………………………….10

3.3.3 High Voltage Operations…………………….……………………11

3.3.4 Patch Panel and Cables……………………………………………12

3.4 Low Voltage System……………………………………...…………13

3.4.1 Low Voltage Power Supply………………………..……………..13

3.4.2 Low Voltage Controls and Operations…………………..………..13

3.4.3 CLAS12 Configuration Change…………………………………..13

3.3.4 Low Voltage Cabling……………………………………………..14

3.5 Light Monitoring System…………………………………...………14

IV.    HTCC authorized personnel…………………………………..…………15

- Introduction

High Threshold Cherenkov Counter (HTCC) is a part of the CLAS12 spectrometer of the Hall B. The detector will be used in experiments with electron beam to reliably detect scattered electrons at high efficiency and generate fast trigger signal. At the same time HTCC will provide low background by rejecting charged -mesons (detection threshold is 4.9 GeV/c). The HTCC is a one unit detector providing efficient acceptance coverage and introducing minimal amount of material.

- Main Components and Equipment

2.1 Containment Vessel

Containment Vessel (CV) contains all major components of the HTCC: Combined Ellipsoidal Mirror, 48 Photomultiplier Tubes with Winston Cones and Magnetic Shields, Fiber Optics for light monitoring. Overall volume is 18.8 m3. The CV is a rigid composite installation of composite Entry Cone with Window and Dish at the upstream side of the detector, composite Exit Window at the downstream side, composite Moller Cup installed around CV axis (beam direction) between Entry and Exit windows. All composite parts are attached to the rigid metal aluminum corps with thin shells shaping the volume. There are 24 small covered ports of rectangular shape cut in side shells (4 ports per sector) sealed with removable relatively thick aluminum foil.  These ports are to be used for maintenance of inside HTCC components such as PMTs, High Voltage dividers, Winston Cones, Compensation Coils of Magnetic Shields, connectors inside. (See picture above and Fig. 1).

PMT Mounting Units

PMT Frame Plate

Exit Window

Entry Window

Entry Cone

Entry Dish

Al Shells

Moller Cup

Target

Fig. 1

2.2 Combined Ellipsoidal Mirror

HTCC Combined Mirror is multifocal ellipsoidal mirror, (see Fig. 2). It is ~8 ft. of diameter and ~0.62 of thickness (less than 135 mg/cm2). There are 12 identical sectors glued together. Combined Mirror assembly tolerance is kept in range of 0.010  – 0.020. Each sector consists of 4 ellipsoidal mirror facets – #1, #2, #3 and #4. Acceptance of one sector covers polar and azimuthal detection angles in ranges of 5° < θ < 35° and Δφ = 30° correspondingly.  There are no overlaps between mirror facets. Gaps (due to glue joints) between adjacent facets are of the order of ~ 0.010 or less.

Δφ = 30°

8 ft.

# 4

# 3

# 2

# 1

Fig. 2

2.3 Photomultiplier Mounting Units

Each HTCC PMT Mounting Unit contains 5″ Photomultiplier Tubes ET- 9823 with quartz face plate (Fig. 4), Winston Cones, (Fig. 5); 3-layer cylindrical concentric Magnetic Shield, (Fig. 3) with Compensation Coil, (Fig. 6).

Case

Magnetic Shield

ET-9823

Fig. 4

Winston Cone

Fig. 5

Coil

Fig. 6

Fig. 3

All 48 PMT Mounting Units are identical and are grouped in 12 identical sets (4 units per set) installed symmetrically around beam direction. Correspondingly currents in the Compensation Coils are different due to fringe field that is azimuthally symmetric around beam direction and depends on polar angle θ. Compensation currents have to be defined for any given configuration of Central Solenoid and Torus. Each PMT Mounting Unit has been aligned along its axis and in the plane perpendicular to the axis.

PMTs in the HTCC utilize high voltage of positive polarity in order to minimize dark noise. Standard HV dividers provide high linearity of signals. The HV dividers have been modified by adding fast preamplifiers integrated into them allowing to have two identical output signals and also to operate PMT at lower high voltage. The preamplifiers do not need separate power supply. Photomultiplier tube is directly coupled to the Winston Cone since only 80% of Cerenkov photons are impinging the photocathodes. PMT with Winston Cone are installed inside of 3-layer cylindrical Magnetic Shield. Solenoidal Compensation Coil is installed on the innermost layer of the metal Magnetic Shield.

On the downstream wall of the Containment Vessel there are two patch panels with 24 signal and 12 high voltage connectors on each panel. On the same wall there are two multi-pin hermetic connectors for multi-wire cables providing low voltage power for Compensation Coils.

2.4 Light Monitoring System

In order to have PMTs calibrated to provide equal gains the HTCC is equipped with Light Monitoring System (LMS), see Fig. 7. Gain equalization is crucial since number of channels involved in the event detection is up to 4.

4

5

3

2

1

Fig. 7

LMS components shown on Fig. 7 are installed in separate box attached to the Containment Vessel. The box is connected in series to the major CV volume and is purged with the same flow of dry gas. LMS main components are:

- Fast Light Source

- Integrating Sphere 4 of diameter

- Adapter

- Fiber bundle Harness

- Reference PMT with YAP crystal

Fast Light Source (1) is computer controlled remotely operated. Corresponding driver generates signal comparable with signals from fast plastic scintillators. The source can provide stable light that can be adjusted in wide range of intensities. LED light source provides input light to the Integrating Sphere (2) that uniformly distributes the light over 48 bundled up fiber optics (4) and additional fiber that illuminates Reference PMT (5). For additional control the Reference PMT is equipped by scintillating YAP crystal. Adapter with 48 fibers is hermetically sealed providing transition of bundled up fiber optics through the downstream wall of the CV. All 48 fibers are of the same length. Each fiber is attached to the corresponding PMT Mounting Unit and illuminates PMT contained in the unit.

2.5 Gas Supply System

The HTCC is continuously purged with dry CO2 gas whether or not the detector is used in any experiment. On Fig. 8 is given a scheme of the HTCC Gas System Controls.

Fig. 8 HTCC Gas System Controls

Parameters of environment inside/outside of the HTCC are continuously monitored online. The HTCC Gas Controls tab (Fig. 9) is also available on the local monitor.

Fig. 9 HTCC Gas Control Tab

The HTCC is equipped with additional supplemental set of components (oil sieve and Submicronic Gas Filter for additional purification of the CO2 gas, Mass Flow Controller to change gas flow operatively during any maintenance if needed, and Magnetic Transmitter to monitor pressure in the CV ) installed right at the input of the detector. To change flow, enter the desired flow in the CO2 Set Pt control field and monitor the flow on the CO2 Flow indicator.

The exhaust of the CV is on the top of downstream side of the detector and is being regulated by set of bubblers. The line is connected to dry Bubbler and then split between two Bubblers with mineral oil controlling and regulating working pressure.

All adjustments of working parameters of the Gas System including routine maintenance must be performed by authorized personal of Hall B working as required by OSP “Testing and running of the High Threshold Cerenkov Counter of the CLAS12 spectrometer in the experimental Hall B”.

- Standard Operations

The section covers operations with HTCC performed by Shift Warkers during experiments.

  - Alarm Handler

Currently the HTCC control software monitors both Gas System parameters and High Voltage and generates corresponding alarms.

    - Gas System Alarms

Monitored quantities for Gas System are:

- Incoming gas flow rate

- Humidity inside the volume of the detector (in ppm)

- Differential pressure in the detector volume

Corresponding Alarm screens are shown on Fig. 10a, 10b and 10c. None of these alarms should be blindly acknowledged as they might indicate dangerous situation that potentially can be dangerous to detector integrity and compromise its performance.

Low input flow of the gas could indicate that the supply CO2 gas tank is nearly empty and needs to be replaced. This condition will most likely trigger “low pressure” alarm.

Acknowledge the alarm, make logbook entry. Call system expert and follow instructions.  Call Hall B Work Coordinator if the CO2 tank is to be replaced.

High input flow/high pressure alarm requires immediate actions since prolonged high pressure within detector volume is dangerous. Call system expert and follow instructions. Make logbook entry.

- (b)                                                      (c)

Fig.10 Alarm screens for CO2 flor rate (a), Absolute humidity in the HTCC volume (b),

and Differential pressure (c)

    - High Voltage Alarm

High Voltages, currents are continuously monitored. Voltages are supposed to stay within a preset range and currents should not exceed set values. In case of HV alarm Shift Workers can only turn OFF the channel and then back ON. If problem persists then turn channel OFF, make logbook entry, call system expert and follow instructions.

  - Gas Supply System

The HTCC is filed with dry CO2 gas at room temperature and pressure. It has to be continuously purged to keep Relative Humidity (RH) inside the CV at level of 3.5% or lower. Due to diffusion of water vapors through many composite parts the detector needs to be purged rates in range of 7.5 L/min to 10 L/min depending on relative humidity in the Hall B area.

The detector is equipped with continuous controls of inside relative humidity, temperature, pressure, gas flow rate. All monitoring information is accessible on-line and also is displayed on the local monitor placed near the detector in the hall. Established slow controls generate corresponding alarms if critical controlled parameters are out of specified ranges. The detector is also equipped with system that automatically prevents any over-pressurizing, and maintains steady differential pressure of about 0.1 of H2O in wide range of atmospheric fluctuations. Shift Workers have to contact Gas System experts in case of alarm and follow instructions.

  - High Voltage System

    - High Voltage Power Supply

Standard 24-channel CAEN SY2604 power supply with 48-channel Patch Panel is used to provide positive high voltage in range of 1,500 V to 2,000 V to the 48 Photomultiplier Tubes used the HTCC detector. Power per channel needed is considerably lower than CPS modules can deliver. PMTs are equipped with standard high voltage dividers designed to provide high linearity of signals at relatively low signal strength. They are modified to have two identical output signals. Turning power supplies ON, running necessary checks of PMT gains using HTCC LMS, establishing high voltages will be done by system experts.

    - High Voltage Controls

High voltage controls are assembled within one screen, accessible under HTCC/HV in Slow Control menu and allows user to monitor and control high voltages for PMTs in real time. Corresponding sample GUI is shown on Fig.11.

Fig.11. HTCC Main HV control screen showing status of all 48 channels with legend.

Side table (right) shows summary of gas system.

Control options include turning ON/OFF the whole detector as shown on Fig. 12:

Fig.12 Clicking on the HV menu allows turning ON/OFF the detector

Individual sectors or channels also can be operated (by clicking corresponding sector or channel):

Fig.13. Clicking on sector number will open sector specific menu

Status of channels is color coded and is present on HTCC HV screen.

Fig.14. HV single channel control screen

The “Expert” mode, accessible form HV menu, allows changing HV settings for individual channels and must be used by HTCC experts only. Additional information is displayed in right lower corner of the HV screen and provides data on gas system status.

    - High Voltage Operations

Nominal high voltages for PMTs are obtained during calibration runs and established values are loaded to the High Voltage crates with specially designed scripts before starting of the runs. In case of the individual PMT malfunction it is possible to turn individual channel OFF using expert GUI:

Fig.15. HTCC Expert GUI

Shift Workers should operate with high voltages only by turning channels OFF and back ON to restore working status. If this simple procedure does not help the channel has to be turned OFF. Call System Expert and follow recommendations. Nominal settings of other parameters should not be changed under normal circumstances. In case of HV crate failure there is a possibility to restore settings from a backup file using HV Menu/Restore setting option. Set of back up files has to be very carefully maintained and used by a HTCC calibration expert.

    - Patch Panel and Cables

Shift Workers may implement any configuration changes (replacement of bad power supplies, patch panel, cables, connectors and such) as necessary only under direct supervision of system experts. Patch panel (see Fig. 16) used in the HTCC High Voltage System is modified due to excessive noise, generated by CAEN power supplies. Set of low pass filters are installed in the high voltage patch panel.

Fig.16. HV Patch panel. Two high density cables (shown in the bottom) from CAEN

power supply are connected to the Patch panel. High voltage is distributed

between 48 independent channels

Problem with filters can lead to the high voltage cut off, so PMT signals will be lost. In case of complicated failure the high voltage applied to the PMT can significantly deviate from nominal value and therefore PMT rate and signal shape can be very different from baseline. In cases of failure (most likely due to malfunctioning of HV power supply or filters in the patch panel, bad/loose connections) the first measure to minimize the consequences will be to turn off high voltage to the channel. Shift Workers have to contact System Experts and make a note in the logbook. Channel will be checked by expert using NIM standalone power supply. Replacement patch panel should be equipped with 48 filters to operate properly. If this option is not available than separate HV power supply has to be used until patch panel is fixed.

  - Low Voltage System

    - Low voltage Power Supply

Low voltage power supplies are used to energize Compensation Coils installed in the PMTs Magnetic Shields to bring attenuated fringe field to zero level. Turning low voltage power supplies ON, running necessary checks of PMT gains using HTCC LMS, establishing compensation currents for given superconductive magnets configuration will be done by system experts.

3.4.2 Low Voltage Controls and Operations

Low voltage control screen is shown on Fig.17. It allows monitoring of currents through the compensation coils. CLAS12 geometry is such that there are only four different values of magnetic field at which PMTs of HTCC are operating. Correspondingly it requires four different values of the compensation current. Screen to control the currents is located under HTCC/LV menu and shows eight channels. Channels 1 through 4 are identical to channels 5 through 8 and control different halves (portions) of the HTCC detector. Values (including sign) of the currents are determined during magnetic shielding calibration runs performed by system expert.

Fig.17 HTCC Low voltage screen allows to control currents of all compensation coils

3.4.3 CLAS12 Configuration Change

Fringe field in location of PMTs in the HTCC almost completely are defined by Central Superconducting Solenoid (CSS). Compensation currents have to be redefined only if CSS current is changed more than 5%. Due to axial symmetry of the HTCC geometry only 4 current values have to be defined and used for given configuration of superconductive magnets. Shift Workers must contact and inform system expert on expected configuration change.

3.4.4 Low Voltage cabling

There are two low voltage power supply modules. These modules have 4 independent outputs (one is shown in Fig.18) and each module serves the left or right halves of the detector.

Fig.18. Connector installed on the back panel of the Low Voltage power supply.

Output provides current for 6 compensation coils connected in series.

Each of the 4 outputs from one of the power supplies is a 2-wire connection that is combined into an 8 wire MS connector input, which is located next to the signal and HV patch panel, for each half of the detector.

  - Light Monitoring System

HTCC Light Monitoring System will be used for calibration of the detector and only during gain equalization runs performed by system experts.  LMS includes reference photomultiplier tube coupled to the YAP crystal (Yttrium Aluminum Perovskite) and is installed in the LMS box on the downstream panel of the HTCC. YAP crystal is an extremely stable and long term source of gamma radiation deposited onto the scintillating crystal. The LMS box is a confined volume and is marked with a “Radioactivity” sign. Periodical checks for leakages are performed by authorized Radiation Control Group personnel. Never make any changes in the LMS box containing YAP crystal. The YAP crystal can be replaced, removed or relocated from the LMS box only by permission of Radiation Control Group (contact David Hamlette, ext. 7219).

- HTCC authorized personnel

Beyond turning on/off the HTCC channels high voltage and monitoring the system parameters, all other operations, repairs are to be carried out by the list of personnel shown in the following table:

| Name | Telephone | E-mail | Area |
| --- | --- | --- | --- |
| Youri Sharabian | (757)269-5829 | youris@jlab.org | HTCC Group Leader |
| Nick Markov | (757)269-5653 | markov@jlab.org | Software HV, LV power supplies Compensation coils |
| Mark McMullen | (757)269-7738 | mcmullen@jlab.org | Gas System monitoring |
| William Phelps | (757)269-7859 (757)567-5546 (cell) | wphelps@jlab.org | Software HTCC calibration Local Gas System Hardware |
| David Riser | (757)269-5629 | dmriser@jlab.org | HTCC calibration Gas leak tests |